Table 1. Adaptive-threshold gait-event accuracy. HS = heel strike; TO = toe off. Miss rate is the percentage of GaitRite events unmatched within 150 ms. Bias = mean IMU-minus-GaitRite timing error.

| Condition | Event | Miss rate (%) | N matched | Bias (ms) | RMSE (ms) | MAE (ms) |
| --- | --- | --- | --- | --- | --- | --- |
| Single-task | HS | 0.0 | 3749 | −0.7 | 26.8 | 16.7 |
|  | TO | 2.8 | 3698 | +3.9 | 32.8 | 22.7 |
| Serial 2s | HS | 0.2 | 3650 | −0.2 | 31.2 | 18.8 |
|  | TO | 3.2 | 3586 | −0.7 | 37.8 | 25.9 |
| Serial 3s | HS | 0.4 | 3858 | +0.1 | 32.4 | 19.8 |
|  | TO | 3.6 | 3760 | 0.0 | 38.7 | 26.6 |
| Serial 7s | HS | 0.4 | 3554 | +0.1 | 32.6 | 20.1 |
|  | TO | 3.8 | 3499 | +1.7 | 39.3 | 27.3 |

Table 2. Pipeline design comparisons. The blocks compare stance detection and integration anchors on their respective shared stride sets, so N differs. All other pipeline steps were held constant. Bias = IMU minus GaitRite.

| Variant | N pairs | RMSE (cm) | Bias (cm) | MAE (cm) | ICC(2,1) |
| --- | --- | --- | --- | --- | --- |
| Stance detection |  |  |  |  |  |
| Bilateral (canonical) | 10,472 | 3.9 | +0.3 | 2.5 | 0.976 |
| Ipsilateral (threshold-only) | 9,708 | 3.8 | −0.3 | 2.4 | 0.976 |
| Gait-event timing anchor |  |  |  |  |  |
| Mid-stance (canonical) | 10,459 | 3.9 | +0.3 | 2.5 | 0.976 |
| GaitRite heel strike | 10,459 | 4.1 | +1.1 | 2.8 | 0.973 |

Table 3. Bilateral ZUPT stride-length accuracy by condition. n = participants; N = matched stride pairs; bias = mean IMU-minus-GaitRite error; LoA = Bland–Altman 95% limits of agreement.

| Condition | n | N pairs | Bias (m) | RMSE (m) | MAE (m) | 95% LoA (m) | ICC(2,1) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Single-task | 32 | 2,544 | +0.007 | 0.040 | 0.027 | [−0.069, +0.084] | 0.973 |
| Serial 2s | 30 | 2,618 | +0.001 | 0.037 | 0.024 | [−0.072, +0.073] | 0.977 |
| Serial 3s | 32 | 2,770 | +0.000 | 0.039 | 0.024 | [−0.077, +0.077] | 0.975 |
| Serial 7s | 32 | 2,540 | +0.003 | 0.039 | 0.025 | [−0.073, +0.078] | 0.971 |
| Pooled | 32 | 10,472 | +0.003 | 0.039 | 0.025 | [−0.073, +0.078] | 0.976 |

Table 4. Spatiotemporal agreement by condition across 10,472 matched strides. GaitRite and IMU values are mean (SD), and bias = IMU minus GaitRite. Stride-length CV is calculated within participant and reported as the cohort mean.

| Parameter | Condition | GaitRite | IMU | Bias | RMSE | ICC(2,1) |
| --- | --- | --- | --- | --- | --- | --- |
| Gait speed (m/s) | Single-task | 1.186 (0.207) | 1.189 (0.215) | +0.002 | 0.070 | 0.945 |
|  | Serial 2s | 1.039 (0.231) | 1.043 (0.237) | +0.004 | 0.064 | 0.963 |
|  | Serial 3s | 0.992 (0.229) | 0.994 (0.233) | +0.002 | 0.062 | 0.964 |
|  | Serial 7s | 0.994 (0.196) | 0.998 (0.202) | +0.004 | 0.063 | 0.949 |
|  | Pooled | 1.051 (0.230) | 1.054 (0.236) | +0.003 | 0.065 | 0.961 |
| Stride time (s) | Single-task | 1.088 (0.097) | 1.093 (0.110) | +0.005 | 0.063 | 0.816 |
|  | Serial 2s | 1.181 (0.160) | 1.177 (0.158) | −0.004 | 0.071 | 0.900 |
|  | Serial 3s | 1.209 (0.168) | 1.207 (0.168) | −0.002 | 0.071 | 0.911 |
|  | Serial 7s | 1.218 (0.165) | 1.216 (0.165) | −0.002 | 0.073 | 0.903 |
|  | Pooled | 1.175 (0.159) | 1.174 (0.160) | −0.000 | 0.070 | 0.905 |
| Cadence (steps/min) | Single-task | 111.2 (9.9) | 110.9 (11.1) | −0.33 | 6.18 | 0.827 |
|  | Serial 2s | 103.3 (12.8) | 103.7 (13.2) | +0.36 | 5.85 | 0.899 |
|  | Serial 3s | 101.1 (13.1) | 101.3 (13.5) | +0.21 | 5.57 | 0.912 |
|  | Serial 7s | 100.2 (12.5) | 100.4 (12.9) | +0.18 | 5.67 | 0.900 |
|  | Pooled | 103.9 (12.9) | 104.0 (13.4) | +0.11 | 5.82 | 0.902 |
| Stride-length CV (%) | Single-task | 2.69 | 3.24 | +0.56 | 1.12 | 0.504 |
|  | Serial 2s | 3.18 | 3.39 | +0.21 | 0.82 | 0.827 |
|  | Serial 3s | 3.34 | 3.89 | +0.55 | 1.74 | 0.518 |
|  | Serial 7s | 3.12 | 3.79 | +0.66 | 1.60 | 0.489 |
|  | Pooled | 3.08 | 3.58 | +0.50 | 1.38 | 0.576 |