Table 2. Pipeline design comparisons. The blocks compare stance detection and integration anchors on their respective shared stride sets, so N differs. All other pipeline steps were held constant. Bias = IMU minus GaitRite.

| Variant | N pairs | RMSE (cm) | Bias (cm) | MAE (cm) | ICC(2,1) |
| --- | --- | --- | --- | --- | --- |
| Stance detection |  |  |  |  |  |
| Bilateral (canonical) | 10,472 | 3.9 | +0.3 | 2.5 | 0.976 |
| Ipsilateral (threshold-only) | 9,708 | 3.8 | −0.3 | 2.4 | 0.976 |
| Gait-event timing anchor |  |  |  |  |  |
| Mid-stance (canonical) | 10,459 | 3.9 | +0.3 | 2.5 | 0.976 |
| GaitRite heel strike | 10,459 | 4.1 | +1.1 | 2.8 | 0.973 |